EVALUATION OF 4 ELISA TEST KITS FOR ANTIBODIES TO CLASSICAL SWINE FEVER

# C.Y. Tee1, B.K. Lim2, H.L. Too3 and P.Y. Choo4

# 1Rhone Ma Malaysia (M) Sdn Bhd, Selangor, Malaysia, 2Vet Food Agro Diagnostic (M) Sdn Bhd, Selangor, Malaysia, 3Merial Asia Pte. Ltd, 4Asia Pacific Special Nutrients (M) Sdn Bhd, Selangor, Malaysia

chiouyan.tee@rhonema.com

Introduction

# Commercial Enzyme Linked Immunosorbent Assay (ELISA) test kits for classical swine fever (CSF) had been widely used as part of the rapid evaluation tools in disease and antigen exposure. However, validations need to be harmonized on an international level for them to be useful for general interpretation1. The objective of this study was to evaluate four commercial ELISA test kits in their assay of antibody to CSF.

Material and Methods

Sera

Six standard sera obtained from EU Reference Laboratory of Hannover University for CSF, i.e. 3 positive sera with different titer level (low-L, medium-M, high-H) from pig experimentally infected with reference CSFV where sera were collected from pigs more than 21 days-post-infection (dpi), 2 positive sera with titres low-L and medium-M from pigs less than 21 dpi, and 1 negative control. All sera were tested with Neutralization peroxidase-linked antibody assay (NPLA) at the Reference Laboratory.

Serology test

All sera were tested for antibodies to CSF concurrently with 4 commercial test kits, ie SERELISA HCV Ab Mono Blocking (S), BioChek CSF-E2 (B), CIVTESTTM suis HC/PPC (C) and IDEXX Herdcheck*CSFV Ab (I). Duplicate tests were done for each serum on respective test kits. Test results were reported as positive, negative or positive/negative as according to the cut-off points indicated by the manufacturers.

Results and Discussion

# The sensitivity of test kit S, B, C and I were found to be 60%, 100%, 70% and 80% respectively to the standard positive sera (Table 1). With the exception of test kit S, the other test kits were able to detect CSF antibodies in sera from re-convalescent pigs (>21 days post infection). Test kit B was the kit that was able to detect all positive samples from early phase of infection (<21 days post infection). False negative results were reported for test kit S and I for positive samples from pigs less than 21 dpi. A false positive result was found with negative sera tested with S test kit.

# The differences in test kit sensitivity performance may be related to a number of factors, which could include the nature of the antigen coated onto the plate which could include glycoprotein2 as well as the detection cut-off points designed to meet specific requirements and purpose of the test kit, i.e. a lower sensitivity for screening infected herds might not suitable to use for detection of antibody in CSF endemic countries or evaluation of vaccination performance, or in converse situation detection of true positives in CSFV eradication programs.Table 1: Qualitative findings on 3 groups of sera with duplicate tests (grey shaded).

| Standard Sera | Test kit |
| --- | --- |
| Group | Level | S | B | C | I |
| CSF >21dpi | L | Positive | Positive | Positive | Positive |
|  | L-A2 | Positive | Positive | Positive | Positive |
|  | M | Negative | Positive | Positive | Positive |
|  | M-A2 | Negative | Positive | Positive | Positive |
|  | H | Positive | Positive | Positive | Positive |
|  | H-A2 | Positive | Positive | Positive | Positive |
| CSF <21dpi | L | Negative | Positive | Pos/Neg | Positive |
|  | L-B2 | Negative | Positive | Pos/Neg | Positive |
|  | M | Positive | Positive | Positive | Negative |
|  | M-B2 | Pos/Neg | Positive | Pos/Neg | Negative |
| Negative control | N | Positive | Negative | Negative | Negative |
|  | N-2 | Pos/Neg | Negative | Negative | Negative |

Conclusion and Recommendation

This study only showed qualitative results using ELISA test kit on experimental CSFV infected animal under the test interpretation instructions from the manufacturers. More studies are required on efficacy of different test kits in detection of antibodies from pig vaccinated with CSF vaccines.

Literature cited

1. Blome S. et al., (2006) Rev. sci. tech. Off. Int. Epiz. 25(3), 1025-1038,

2. Langedijk et al., (2001) J. Clin Microbiol. 39(3) 906-912.

SERELISA® HCV Ab Mono Blocking is a registered trademark of Synbiotic Corp; BioChek CSF-E2 is a registered trademark of BioChek; CIVTESTTM is a registered trademark of HIPRA Laboratories; IDEXX ®Herdcheck CSF is a registered trademark of IDEXX Laboratories Inc.